## Slide 1: PLUS Toolkit: Acquisition, Processing, Calibration
Kyle Sunderland, Andras Lasso
(1krs1@queensu.ca, lasso@queensu.ca)
May 4, 2021
Laboratory for Percutaneous Surgery
Queen’s University

## Slide 2: Introduction
PLUS – Public software Library for UltraSound imaging
Goal: Facilitate rapid prototyping of ultrasound-guided intervention systems for translational clinical research
Acquire, process, transfer synchronized image and position data
Probe
Image
Needle
Reference
Tracker
Phantom model
Ultrasound
computer
3D Slicer
computer
Notes: Open source software library developed in Perk Lab
Rapidly prototype image guided systems for use in clinical research
Aquire and manage data from imaging and tracking modalities
Merge to single data source

## Slide 3: Introduction
My Algorithm
My Interface
My Tracker
My Algorithm v2
My Interface v2
My Algorithm
PLUS
My Tracker
My New Tracker
My New Tracker
Notes: To use algorithm + tracker software might be implemented like. (A)
Good for 1, but for >1 results in duplicated work
Need to implement new interface and modify algorithm to support
Plus: only need to adjust single file to completely change hardware
Output from Plus always in consistent format
Many devices supported
Adding support only requires interface between Plus and hardware

## Slide 4: Plus Server Application
Connect to hardware devices
Broadcast data via OpenIGTLink (e.g. to 3D Slicer)
Remote control access from 3D Slicer
  Start/stop device connections
  Data recording
  Volume reconstruction
Notes: Main application: PlusServer
Connect devices and transmit data over OpenIGTLink
(Open source network communication protocol)
Plus server launcher
  Select configuration and start
Controlled remotely from Slicer

## Slide 5: Plus Server Application
Probe
Image
Needle
Reference
Tracker
Phantom model
PLUS
Slicer
Ultrasound
computer
3D Slicer
computer
OpenIGTLink
Notes: Single config file, set up connections like…
Describe connections

## Slide 6: System Overview
Hardware (ultrasound, tracker, etc.)
PLUS
SlicerRT
Slicer
…
PerkTutor
SlicerIGT (Calibration, Visualization)
Image Overlay
…
Notes: End up with workflow that looks like this:
Low level hardware
Connected to Plus
Broadcast to Slicer over OpenIGTLink
Processed in modules

## Slide 7: Data Acquisition - Tracking
Ascension EM tracker
NDI Aurora, Polaris, and Certus optical and electromagnetic trackers
Claron MicronTracker optical tracker
OptiTrack optical trackers
Brachy steppers (CMS Accuseed, Burdette Medical systems, CIVCO)
PhidgetSpatial inertial measurement device
CHRobotics inertial measurement device
3dConnexion SpaceNavigator 3D mouse
OpenIGTLink (for BrainLab, Siemens MRI scanners, and other compatible devices)
Software devices: file source, US simulator
Notes: Plus supports many different tracking devices (list is not exhaustive)
Red stars are hardware that we have access to
Green stars are virtual devices that don’t require a piece of hardware

## Slide 8: Data Acquisition - Imaging
Ultrasonix: B-mode & RF (through research interface)
BK ProFocus: B-mode & RF (through research interface)
Telemed ultrasound
Clarius ultrasound
ImagingControl framegrabbers
Epiphan framegrabbers
Webcam video
OpenIGTLink (for MUSiiC, Siemens MRI scanners, and other compatible devices)
Other software devices: file source, US simulator
Notes: Similarily, Plus also supports a wide variety of Imaging devices

## Slide 9: fCal Application
Spatial ultrasound calibration
Temporal ultrasound calibration
fCal phantom
fCal (Freehand tracked ultrasound calibration application)
Notes: Other application fCal
- Used for temporal + spatial calibration
Temporal calibration
  How
  Why
Spatial calibration
  How

## Slide 10: Spatial Calibration
Goal: determine the transform between coordinates systems
Notes: Spatial calibration
  Why
- Corner of the image relative to the probe
- Scale of the image in mm

## Slide 11: Sequence Files
Timestamped image and transform recording
Based on the Meta IO standard image file format (.mha, .mhd)
Images readable by many existing applications (e.g. 3D Slicer)
Extra information for tracking/reconstruction
  Image orientation
  Transforms from tracker
  Tracker and image status
  Timestamps
  etc.
Usable by all Plus applications
Notes: In addition to streaming data, Plus can also record data in Sequence files
Contain synchronized image and transform data
Based on the MetaIO format developed in ITK
Sequences can be used in other applications such as 3D Slicer
Both PlusServer and fCal can read and replay sequence files, as well as perform other algorithms such as volume reconstruction

## Slide 12: Sequence Files
EditSequenceFile.exe
  Editor for sequence files
ViewSeqMetaFile.exe
  Replay saved sequence files
ObjectType = Image
NDims = 3
BinaryData = True
BinaryDataByteOrderMSB = False
CompressedData = True
DimSize = 820 616 97
ElementType = MET_UCHAR
…
UltrasoundImageOrientation = MF
Seq_Frame0000_ProbeToTrackerTransform = 0.956683 -0.263308 0.124204 -190.886 0.269031 0.962616 -0.0315089 -98.0911 -0.111264 0.0635588 0.991756 -1949.07 0 0 0 1
Seq_Frame0000_ProbeToTrackerTransformStatus = OK
Seq_Frame0000_ReferenceToTrackerTransform = 0.260287 0.964665 0.0408926 -208.632 -0.96117 0.262899 -0.0838866 -115.511 -0.0916731 -0.0174702 0.995636 -1992.7 0 0 0 1
Seq_Frame0000_ReferenceToTrackerTransformStatus = OK
Seq_Frame0000_StylusToTrackerTransform = 1 -8.43482e-005 0.000190641 0.370779 8.39344e-005 1 0.000146179 0.316166 -0.00019069 -0.000146138 1 -0.0640528 0 0 0 1
Seq_Frame0000_StylusToTrackerTransformStatus = INVALID
Seq_Frame0000_Timestamp = 345.627957
Seq_Frame0000_ImageStatus = OK
…
Seq_Frame0096_StylusToTrackerTransform = 1 0.000362791 -8.01739e-005 -0.106056 -0.000363818 1 -0.000369765 -0.839888 8.00139e-005 0.000370927 1 0.0862409 0 0 0 1
Seq_Frame0096_StylusToTrackerTransformStatus = INVALID
Seq_Frame0096_Timestamp = 355.783014
Seq_Frame0096_ImageStatus = OK
ElementDataFile = LOCAL
xœìÝw...
Notes: Point out features
  Metadata
  Transform
  Status
Other applications:
  Edit
  Replay

## Slide 13: PLUS Configuration Files
Notes: Need to be able to write config files
What are they?
  XML
  What hardware
  What processing
  What connections (devices)
  What output

## Slide 14: XML Basics
Element: XML element is an entry in an XML file
<Unicorn/>
Attribute: XML attribute is a property of an element
<Unicorn realness=“notReal” size=“12” />
Example:
<Animals>
<Animal species=“Unicorn” realness=“notReal” size=“12”/>
<Animal species=“Elephant” realness=“real” size=“50”/>
</Animals>
Notes: Element
Object in tag
Has name
Attribute:
Defines properties of an element
Can have nested elements.

## Slide 15: Config File Terminology
Hardware: Any physical or virtual device
An EM position tracker
Device: The interface to a hardware
Interface to EM position tracker
Data Source: A construct around a buffer containing meta information (e.g. what type of source)
The tracking data for the ultrasound probe
Channel: A construct representing a stream of data
The combined tracking data

## Slide 16: Config File Terminology
Transform: A 4x4 matrix representing a displacement in 3D (translation, rotation, scaling)
Tracked Frame: An image with n associated transforms (merged from multiple data sources)

## Slide 17
Pipeline
EM Tracker
Ultrasound
Image Channel
Tracking Channel
Virtual Mixer
Virtual Capture
Data Flow
Combined Channel
Visualization Client
Devices are the creation point of data
  Record, simulate, process “external” data
Channels are the communication point between devices
  Outputs tracked frames for client consumption
Notes: Explain what a mixer device is
Explain what a capture device is
Example pipeline in Plus
  Inputs are:
  Produce data channels
Virtual mixer
  Combine channels to a single output
Virtual capture
  Record data
Output to client

## Slide 18
TrackerDevice
Example: Tracked Video
Internal data sources
Output Channel
ProbeToTracker
ReferenceToTracker
StylusToTracker
ProbeToTracker
ReferenceToTracker
StylusToTracker
MixerDevice
Output Channel
ProbeToTracker
ReferenceToTracker
StylusToTracker
Input Channel
ProbeToTracker
ReferenceToTracker
StylusToTracker
VideoDevice
Internal data sources
Output Channel
Video
Video
Input Channel
Video
Video
Notes: By setting up the XML in this format, we have specified
What we are retrieving from hardware (transform, image names)
What devices we are transmitting the information to
What data we are sending over OpenIGTLink

## Slide 19
<Device Id=“MixerDevice" Type="VirtualMixer" >
<InputChannels>
<InputChannel Id="TrackerStream" />
<InputChannel Id="VideoStream" />
</InputChannels>
<OutputChannels>
<OutputChannel Id="TrackedVideoStream"/>
</OutputChannels>
</Device>
Example: Tracked Video
Linking input & output via channels
  Device -> channel -> device (repeat)
<Device
Id="TrackerDevice“
Type="Ascension3DG“
ToolReferenceFrame="Tracker"
... >
<DataSources>
<DataSource Type="Tool" Id="Probe" PortName="0" />
<DataSource Type="Tool" Id="Reference" PortName="1" />
<DataSource Type="Tool" Id="Stylus" PortName="2" />
</DataSources>
<OutputChannels>
<OutputChannel Id="TrackerStream" >
<DataSource Id="Probe"/>
<DataSource Id="Reference"/>
<DataSource Id="Stylus"/>
</OutputChannel>
</OutputChannels>
</Device>
<Device
Id="VideoDevice“
Type="SonixVideo“
ImageToTransducerTransformName="ImageToTransducer“
... >
<DataSources>
<DataSource
Type="Video“
Id="Video“
PortName="B“
PortUsImageOrientation="UF" />
</DataSources>
<OutputChannels>
<OutputChannel
Id="VideoStream“
VideoDataSourceId="Video" />
</OutputChannels>
</Device>
Notes: Emphasize the linkage by ID, give them a basis to understand other config files
Example pipeline
using an Ascension EM tracker and an Ultrasonix ultrasound machine

## Slide 20
Configuration Files
<PlusConfiguration> - Root element
<DeviceSet/> - Information about the config file
<DataCollection> - Where the devices/channels all go
<Device/> - Physical or virtual hardware
</ DataCollection>
<CoordinateDefinitions/> - Constant transforms
<PlusOpenIGTLinkServer/> - What to send over OpenIGTLink
</PlusConfiguration>
Notes: Emphasize the linkage by ID, give them a basis to understand other config files
Within the config file, the elements can be stored within these tags
DataCollection contains all of the device and channel information
And finally CoordinateDefinitions can contain hardcoded transforms (for applications like ultrasound calibrations)
PlusOpenIGTLinkServer sends data from a single channel over openigtlink using the specified connection information

## Slide 21: Some Advice
Do NOT start from scratch
Plus website contains documentation and config files for devices: http://perk-software.cs.queensu.ca/plus/doc/nightly/user/Configuration.html

## Slide 22: PLUS Configuration File Tutorial

## Slide 23: Laptops/Webcams/Markers
Plus is only available to download for Windows
Hands-on tutorial requires a Windows computer + webcam
Have your printed ArUco markers ready

## Slide 24: Download & Install PLUS
From: http://perk-software.cs.queensu.ca/plus/packages/stable/
  Choose: PlusApp-…-Win64.exe
  NOT Ultrasonix/Telemed/etc.

## Slide 25: Download & Install Slicer
From: https://download.slicer.org/
  Choose: Slicer 4.11.20210226 (Stable release)
Install SlicerOpenIGTLink extension
  Open Extensions manager
  Switch to “Install Extensions” and search for “SlicerOpenIGTLink”
  Click INSTALL
  Restart Slicer

## Slide 26: PLUS Configuration File Tutorial: Webcam

## Slide 27: Start PLUS Server Launcher

## Slide 28: Select Configuration File Directory
1
3
Directory where you installed PLUS (C:/Users/<YourUserName>/PlusApp-…-Win32/config)
Choose “config” subdirectory
2

## Slide 29: Choose Configuration for Color Video
1
2

## Slide 30: Launch Server
May take a few seconds…
1

## Slide 31: Check Connection
See

## Slide 32: Check Message Log
1

## Slide 33: Connect from Slicer
1
2

## Slide 34: Connect from Slicer
1
2
Should populate with
Scene > IGTLConnector > IN > Image_Reference
See

## Slide 35: View Sent Images
2
1

## Slide 36: View Sent Images
1
See

## Slide 37: Edit Configuration Files
1
See

## Slide 38: Orienting Image Properly
<Device
Id="VideoDevice"
Type="MmfVideo"
FrameSize="640 480"
VideoFormat="YUY2"
CaptureDeviceId="0" >
<DataSources>
<DataSource Type="Video" Id="Video" PortUsImageOrientation="UN" ImageType="RGB_COLOR" />
</DataSources>
<OutputChannels>
<OutputChannel Id="VideoStream" VideoDataSourceId="Video" />
</OutputChannels>
</Device>
Change to “UN”
(if this does not work, try “MN”)
2
1,3

## Slide 39: View Sent Images
See

## Slide 40: Increase Image Resolution
<Device
Id="VideoDevice"
Type="MmfVideo"
FrameSize=“1280 720"
VideoFormat="YUY2"
CaptureDeviceId="0" >
<DataSources>
<DataSource Type="Video" Id="Video" PortUsImageOrientation="UN" ImageType="RGB_COLOR" />
</DataSources>
<OutputChannels>
<OutputChannel Id="VideoStream" VideoDataSourceId="Video" />
</OutputChannels>
</Device>
Change to “1280 720”
Not all webcams support all resolutions (check for list of supported resolutions in PLUS message log)
2
1,3

## Slide 41: View Sent Images
See
Re-center

## Slide 42: Crop Image
<Device
Id="VideoDevice"
Type="MmfVideo"
FrameSize=“1280 720"
VideoFormat="YUY2"
CaptureDeviceId="0" >
<DataSources>
<DataSource Type="Video" Id="Video" PortUsImageOrientation="UN“ ImageType="RGB_COLOR" ClipRectangleOrigin="300 300" ClipRectangleSize=“400 400" />
</DataSources>
<OutputChannels>
<OutputChannel Id="VideoStream" VideoDataSourceId="Video" />
</OutputChannels>
</Device>
Add ClipRectangleOrigin="300 300" ClipRectangleSize=“400 400"
2
1,3

## Slide 43: View Sent Images
See
Re-center

## Slide 44: PLUS Configuration File Tutorial: Optical Marker Tracking

## Slide 45: Start PLUS Server Launcher

## Slide 46: Select Configuration File Directory
1
3
Directory where you installed PLUS (C:/Users/<YourUserName>/PlusApp-…-Win32/config)
Choose “config” subdirectory
2

## Slide 47: Choose Configuration for Optical Tracking
1
2

## Slide 48: Edit Configuration Files
1
See

## Slide 49: Remove Unused Markers
<Device
Id="TrackerDevice"
Type="OpticalMarkerTracker"
CameraCalibrationFile="OpticalMarkerTracker/realsense_gen2_calibration.yml"
ToolReferenceFrame="Tracker"
TrackingMethod="OPTICAL"
MarkerDictionary="ARUCO_MIP_36h12">
<DataSources>
<DataSource Type="Tool" Id="Marker0" MarkerId="0" MarkerSizeMm="80" />
<DataSource Type="Tool" Id="Marker1" MarkerId="1" MarkerSizeMm="80" />
</DataSources>
<InputChannels>
<InputChannel Id="VideoStream" />
</InputChannels>
<OutputChannels>
<OutputChannel Id="TrackerStream" >
<DataSource Id="Marker0"/>
<DataSource Id="Marker1"/>
</OutputChannel>
</OutputChannels>
</Device>
We will only use two markers for our setup.
We use markers 0 and 1.
Remove here
Remove here

## Slide 50: Change Marker Sizes
<Device
Id="TrackerDevice"
Type="OpticalMarkerTracker"
CameraCalibrationFile="OpticalMarkerTracker/realsense_gen2_calibration.yml"
ToolReferenceFrame="Tracker"
TrackingMethod="OPTICAL"
MarkerDictionary="ARUCO_MIP_36h12">
<DataSources>
<DataSource Type="Tool" Id="Marker0" MarkerId="0" MarkerSizeMm=“40" />
<DataSource Type="Tool" Id="Marker1" MarkerId="1" MarkerSizeMm=“40" />
</DataSources>
<InputChannels>
<InputChannel Id="VideoStream" />
</InputChannels>
<OutputChannels>
<OutputChannel Id="TrackerStream" >
<DataSource Id="Marker0"/>
<DataSource Id="Marker1"/>
</OutputChannel>
</OutputChannels>
</Device>
Change to “40” here.
Our printed markers are 40mm x 40 mm.

## Slide 51: Change Sent Transforms
<PlusOpenIGTLinkServer
MaxNumberOfIgtlMessagesToSend="1"
MaxTimeSpentWithProcessingMs="50"
ListeningPort="18944"
SendValidTransformsOnly="true"
OutputChannelId="TrackerStream" >
<DefaultClientInfo>
<MessageTypes>
<Message Type="TRANSFORM" />
</MessageTypes>
<TransformNames>
<Transform Name="Marker1ToMarker0" />
<Transform Name="Marker0ToTracker" />
<Transform Name="Marker1ToTracker" />
</TransformNames>
</DefaultClientInfo>
</PlusOpenIGTLinkServer>
We only have transforms from markers 0 and 1.
Remove here
Add the transform “Marker1ToMarker0”.
Add here

## Slide 52: Launch Server
May take a few seconds…
1

## Slide 53: Check Connection
See

## Slide 54: Check Message Log
1

## Slide 55: Connect from Slicer
1
2

## Slide 56: Connect from Slicer
1
3
Should populate with
Scene > IGTLConnector_1 > IN > Image_Image
See
Change to 18945.
2

## Slide 57: Connect from Slicer
1
2
Should populate with
Scene > IGTLConnector > IN > Marker0ToTracker, Marker1ToTracker, Marker0ToMarker1
See

## Slide 58: View Sent Images
2
1

## Slide 59: View Sent Images
1
See

## Slide 60: View Sent Transforms
See
3
2
1
4

## Slide 61: PLUS Remote Control From Slicer

## Slide 62: Open Config File Directory
1

## Slide 63: Add Config File to Slicer
2
1

## Slide 64: View Config File
1
2

## Slide 65: View Config File
See
1

## Slide 66: Open PlusRemote Module
1
2

## Slide 67: Create PlusServerLauncher
1

## Slide 68: Add Server
1
3
2

## Slide 69: Start Server
1

## Slide 70: Check Plus
2
1

## Slide 71: Save with Scene
1
2
3

## Slide 72: Using plus.xml Extension
1

## Slide 73: Acknowledgements
PLUS (www.plustoolkit.org)
Andras Lasso, Tamas Heffter, Adam Rankin, Csaba Pinter, Tamas Ungi, and Gabor Fichtinger, "PLUS: Open-source toolkit for ultrasound-guided intervention systems", IEEE Trans Biomed Eng. 2014 Oct;61(10):2527-37.
Kyle Sunderland (1krs1@queensu.ca)
Andras Lasso (lasso@queensu.ca)