## Slide 1: Prototyping image-guided therapy applications
David García Mato, Ph.D.
R&D engineer at Ebatinca S.L.
Las Palmas de Gran Canaria, Spain
Mónica García Sevilla, Ph.D.
Universidad de las Palmas de Gran Canaria
Las Palmas de Gran Canaria, Spain

## Slide 2: Outline
Introduction to image-guided therapy (IGT)
Tracking systems
Coordinate systems
IGT setup
Tool calibration
Image-to-patient registration
Examples of IGT applications in Slicer
Virtual and Augmented Reality applications

## Slide 3: Introduction to IGT

## Slide 4: Image-Guided Therapy
Use medical imaging to plan, perform and evaluate surgical procedures
Objective  less invasive and more precise surgeries  shorter hospital stays
Common imaging modalities: CT and MRI. Others: US, angiography…
Use of special devices and software for tracking

## Slide 5: Example
Intense headache. Painkillers not effective…
Tumor in the brain?
Medical Imaging
for Diagnosis

## Slide 6: Example
Benign or malignant?

## Slide 7: Example
Take patient to the OR  Perform the biopsy

## Slide 8: Example
?
Where to insert the needle?
Which trajectory to follow?
When has the needle reached the tumor?

## Slide 9: Example: IGT
Where to insert the needle?
Which trajectory to follow?
When has the needle reached the tumor?

## Slide 10: How to perform this guidance or navigation?
Notes: And how does it work? Well, the idea is very similar to the navigation systems we are used to for the phone or the car

## Slide 11: How to perform this guidance or navigation?
GPS NAVIGATION
SURGICAL NAVIGATION
Map + Position + Path
Medical Image + Tool Position + Path

## Slide 12: How to perform this guidance or navigation?
Map
Medical Images
Tracking system
We navigate by knowing our position in our map
How do we get our position?
Latitude
Longitude
3D coordinates space

## Slide 13: Tracking Systems

## Slide 14: Tracking systems

## Slide 15: Mechanical tracking systems
Description:
Object to be tracked at the end of an articulated arm
Several sensors measure the angles between arm segments
Knowledge of segment sizes and angles provides the position and orientation of the tool
6 degrees of freedom
System is entirely self-referential:
rigid fixation of both the patient and the navigation arm is an important prerequisite

## Slide 16: Mechanical tracking systems
Da Vinci Surgical System (Intuitive)

## Slide 17: Mechanical tracking systems
Patient cart
Console
Vision cart

## Slide 18: Mechanical tracking systems
ADVANTAGES
DISADVANTAGES
High precision
High update rate
Low susceptibility to failure
Impractical handling during some surgeries
Restricted range and mobility
Space requirements
Usually only one instrument can be tracked

## Slide 19: Optical tracking systems (OTS)
IR OTS
Visible light OTS

## Slide 20: Optical tracking systems: IR
Cam. 1
Cam. 2
X
Y
Z
X
Y
Z
Active marker spheres
Passive marker spheres

## Slide 21: Optical tracking systems:IR
Cam. 1
Cam. 2
X
Y
Z
X
Y
Z
Active marker spheres
Passive marker spheres

## Slide 22: Optical tracking systems:IR
Camera A
Camera B
Point detected in every camera defines line in 3D
Line intersection provides 3D position
Known Spatial relationship between cameras
X
Y
Z
TRIANGULATION

## Slide 23: Optical tracking systems: IR
https://www.youtube.com/watch?v=D-jCxeWXt2o&ab_channel=MemoryTree

## Slide 24: Optical tracking systems: IR

## Slide 25: Optical tracking systems: IR

## Slide 26: Optical tracking systems: IR
Tools with passive spherical markers:
Cameras:
Passive spherical markers:

## Slide 27: Optical tracking systems: IR
Optitrack:
Lower precision
Cheaper

## Slide 28: Optical tracking systems: Videometric
Description:
Standard cameras
Visible light
Geometrical patterns
Claron
MicronTracker

## Slide 29: Optical tracking systems: Videometric
Claron MicronTracker
http://www.clarontech.com/microntracker-specifications.php
Specs:
Measurement rate 16-20 Hz
Accuracy 0.2-0.35 mm
Lag ~60 ms
Concurrent markers: 100
Illumination range: 50-100,000 lux

## Slide 30: Optical tracking systems: Videometric
Claron MicronTracker
https://www.youtube.com/watch?v=lfZeXabDjMg&t=19s

## Slide 31: Optical tracking systems
ADVANTAGES
DISADVANTAGES
Accuracy
Reliability
Working volume
Line-of-sight required
Tool design
Tool use in real environment

## Slide 32: Optical vs Electromagnetic tracking systems
Marker
Sensors
IR light
Electromagnetic field

## Slide 33: Electromagnetic tracking systems (EMTS)
NDI Aurora
NDI 3D Guidance
for needles

## Slide 34: Electromagnetic tracking systems (EMTS)

## Slide 35: Electromagnetic tracking systems (EMTS)
Serial port connection
Tracked Tool (with sensor inside)
SCU
(System Control Unit)
Field generator
To PC
To power source
Power connection

## Slide 36: Electromagnetic tracking systems (EMTS)

## Slide 37: Electromagnetic tracking systems (EMTS)
ADVANTAGES
DISADVANTAGES
Line-of-sight not required
Tool tip tracking (small sensors)
Ferromagnetic distortions
Working volume
Limited accuracy

## Slide 38: Coordinate Systems

## Slide 39: Coordinate system: definition
Left-handed or right-handed
Define origin:
  Where is (0,0,0)?
Define directions
  Which direction is X, Y, and Z?
Define units
  How much is 1?
X
Y
Z

## Slide 40: Coordinate system
X
Y
Z
Trackers provide location and/or position information relative to a coordinate system.
(x,y,z) (rx,ry,rz)
(0,0,0)
Tracker coordinate system
Tool coordinate system
ToolToTrackerTransform
ROTATION
TRANSLATION
(0,0,0)

## Slide 41: Transforms: examples
"Identity transform"
Does nothing
"Translation"
Pushes away
"Rotation"
"Scaling"
(and translation)

## Slide 42: Coordinate system
X
Y
Z
Tracker
Tool1-To-Tracker
X
Y
Z
Tool 1
Tool2-To-Tracker
X
Y
Z
Tool 2
Tracker-To-Tool2
Tool1-To-Tool2
[Tracker-To-Tool2] = [Tool2-To-Tracker] -1
[Tool1-To-Tool2] = [Tool1-To-Tracker] * [Tracker-To-Tool2]
Transforms
describe every linear spatial relationship
internally represented by 4x4 matrices
concatenated by matrix multiplication
Tool1-To-Tracker
Tracker-To-Tool2

## Slide 43: Where do transforms come from?
Tracking system
Registration/Calibration
Transform is updated to provide the “real-time” position of a given object
Transform stays constant until next registration/calibration

## Slide 44: Example of IGT setup

## Slide 45: Navigate needle insertion using optical tracker
https://www.youtube.com/watch?v=d_vHpBRZJco

## Slide 46: Setup
Tracker
Phantom
Reference
Needle Tip
Needle
Needle-To-Tracker
Tracker-To-Reference
NeedleTip-To-Needle
Reference-To-Phantom

## Slide 47: Tool tracking
Needle-To-Tracker
Reference-To-Tracker

## Slide 48: Tool tracking
Tracking System
PLUS Toolkit
Tracker SDK
Imaging SDK
… SDK
Synchronization
Calibration
OpenIGTLink
3D Slicer
SlicerIGT
Registration
OpenIGTLink
Visualization
Segmentation
Quantification

## Slide 49: Setup
Tracker
Phantom
Reference
Needle Tip
Needle
Needle-To-Tracker
Tracker-To-Reference
NeedleTip-To-Needle
Reference-To-Phantom

## Slide 50: Setup
Tracker
Phantom
Reference
Needle Tip
Needle
Needle-To-Tracker
Tracker-To-Reference
NeedleTip-To-Needle
Reference-To-Phantom

## Slide 51: Needle Calibration
Needle
?
?
?
?
Needle Tip

## Slide 52: Needle Calibration
Needle Tip
Needle
X
Y
Z
X
Y
Z
NeedleTip-To-Needle

## Slide 53: Needle Calibration
Needle Tip
Pivoting

## Slide 54: Tool calibration: how to compute it?
Pivot calibration can be easily done using the “Pivot Calibration” module within SlicerIGT extension.
ROTATION
TRANSLATION

## Slide 55: Setup
Tracker
Phantom
Reference
Needle Tip
Needle
Needle-To-Tracker
Tracker-To-Reference
NeedleTip-To-Needle
Reference-To-Phantom

## Slide 56: Setup
Tracker
Phantom
Reference
Needle Tip
Needle
Needle-To-Tracker
Tracker-To-Reference
NeedleTip-To-Needle
Reference-To-Phantom

## Slide 57: Phantom Registration
Reference
Phantom
X
Y
Z
X
Y
Z
Reference-To-Phantom

## Slide 58: Phantom Registration
Phantom landmarks
(recorded directly over the 3D model/volume of the phantom)

## Slide 59: Phantom Registration
Reference landmarks
(recorded using tracked tool relative to “Reference” coordinate system)
1
2
3
4
5

## Slide 60: Phantom Registration: computation
1
2
3
4
5
1
2
3
4
5
Phantom Fiducials
(relative to phantom coordinate system)
Reference Fiducials
(relative to reference coordinate system)

## Slide 61: Phantom Registration: computation
1
2
3
4
5
1
2
3
4
5
Phantom Fiducials
(relative to phantom coordinate system)
Reference Fiducials
(relative to reference coordinate system)
Reference-To-Phantom
Minimize distance between homologous points

## Slide 62: Phantom Registration: how to compute it?
Landmark registration can be easily done using the “Fiducial registration wizard” module within SlicerIGT extension.
ROTATION
TRANSLATION

## Slide 63: Setup
Tracker
Phantom
Reference
Needle Tip
Needle
Needle-To-Tracker
Tracker-To-Reference
NeedleTip-To-Needle
Reference-To-Phantom

## Slide 64: Setup
Reference-To-Phantom
Tracker
Phantom
Reference
Needle Tip
Needle
Needle-To-Tracker
Tracker-To-Reference
NeedleTip-To-Needle

## Slide 65: Navigate needle insertion using optical tracker
https://www.youtube.com/watch?v=d_vHpBRZJco

## Slide 66: Examples of IGT applications
Computer-assisted breast tumor excision (LumpNav): https://www.youtube.com/watch?v=gSz8IHmogMo
Computer-assisted orbital floor reconstruction: https://www.youtube.com/watch?v=OzqBqOZsc7M
Computer-assisted dental implant placement: https://www.youtube.com/watch?v=ytq9V2paDzc
Computer-assisted craniofacial surgery (CranioNav): https://doi.org/10.1038/s41598-019-54148-4
Training needle insertion skills (PerkTutor): https://www.youtube.com/watch?v=KfdQ3WwTQHw&t=27s
Forceps-assisted delivery training: https://www.youtube.com/watch?v=EEasWbH1jZI
SlicerVR: https://www.youtube.com/watch?v=F_UBoE4FaoY
Collaborative virtual reality: https://www.youtube.com/watch?v=Sw3JyKfvW6Q

## Slide 67: Prototyping image-guided therapy applications
David García Mato, Ph.D.
R&D engineer at Ebatinca S.L.
Las Palmas de Gran Canaria, Spain
Mónica García Sevilla, Ph.D.
Universidad de las Palmas de Gran Canaria
Las Palmas de Gran Canaria, Spain